## Comparison_table
Spec Group | Spec Code | Specification | Required Value | PolarMed GbhM | CryNova Ltd | ArcticTech Pvt Ltd | MedFreez Corp | BioNordic AS | PolarMed GbhM Compliance | CryNova Ltd Compliance | ArcticTech Pvt Ltd Compliance | MedFreez Corp Compliance | BioNordic AS Compliance | PolarMed GbhM Compliance Score | CryNova Ltd Compliance Score | ArcticTech Pvt Ltd Compliance Score | MedFreez Corp Compliance Score | BioNordic AS Compliance Score
PERFORMANCE | PER-01 | Temperature Range | +2 to +8 °C | +2 to +8 °C | +3 to +9 | +2 to +10 °C | +2 to +8 °C | +1 to +7 °C | Compliant | Partially Compliant | Partially Compliant | Compliant | Partially Compliant | 1 | 0.5 | 0.5 | 1 | 0.5
PERFORMANCE | PER-02 | Temp Fluctuation | ±1 °C | ±1 °C | ±2 °C | ±2 °C | ±2 °C | ±2 °C | Compliant | Not Compliant | Not Compliant | Not Compliant | Not Compliant | 1 | 0 | 0 | 0 | 0
POWER | POW-01 | Power Supply | 220-240 V AC | 230 V AC / 50 Hz / Type C | 230 V AC / 60 Hz / Type C | 220 V AC / 50 Hz / Type G | 120 V AC / 60 Hz / Type A | 230 V AC / 50 Hz / Type F | Compliant | Partially Compliant | Partially Compliant | Not Compliant | Partially Compliant | 1 | 0.5 | 0.5 | 0 | 0.5
POWER | POW-02 | Power Consumption | ≤ 0.9 kWh/day | 1.1 kWh/24h | 1.3 kWh/24h | 1.4 kWh/24h | 1 kWh/24h | 0.9 kWh/24h | Not Compliant | Not Compliant | Not Compliant | Not Compliant | Compliant | 0 | 0 | 0 | 0 | 1
SAFETY | SAF-01 | Alarm Type | Visual + Audible | Visual + Audio | Visual + Audio | Visual + Audio | Visual + Audio | Visual + Audio | Compliant | Compliant | Compliant | Compliant | Compliant | 1 | 1 | 1 | 1 | 1
SAFETY | SAF-02 | Battery Backup Duration | ≥ 2 hours | 12 Hours | 8 Hours | 8 Hours | 8 Hours | 8 Hours | Compliant | Compliant | Compliant | Compliant | Compliant | 1 | 1 | 1 | 1 | 1
DESIGN | DES-01 | Gross Internal Capacity | ≥ 150 L | 150 Liters | 150 Liters | 160 Liters | 155 Liters | 145 Liters | Compliant | Compliant | Compliant | Compliant | Not Compliant | 1 | 1 | 1 | 1 | 0
DESIGN | DES-02 | Internal Shelves | ≥ 3 | 4 adjustable, SS/plastic | 3 adjustable, SS/plastic | 4 adjustable, SS/plastic | 4 adjustable, SS/plastic | 4 adjustable, SS/plastic | Compliant | Compliant | Compliant | Compliant | Compliant | 1 | 1 | 1 | 1 | 1
DESIGN | DES-03 | External Size (W×D×H) | 600×600×1800 mm | ≤ 600×580×1750 mm | 610×590×1770 mm | 590×600×1750 mm | 580×580×1780 mm | 590×590×1750 mm | Compliant | Partially Compliant | Compliant | Compliant | Compliant | 1 | 0.5 | 1 | 1 | 1
CERTIFICATION | CER-01 | CE Marking | Required | CE | CE | — | CE | CE | Compliant | Compliant | Not Compliant | Compliant | Compliant | 1 | 1 | 0 | 1 | 1
CERTIFICATION | CER-02 | ISO 13485 | Required | ISO 13485 | — | ISO 9001 | FDA | ISO 13485, RoHS | Compliant | Not Compliant | Not Compliant | Not Compliant | Compliant | 1 | 0 | 0 | 0 | 1
LOGISTICS | LOG-01 | Delivery Timeline | ≤ 35 days | 28 Days | 40 Days | 35 Days | 32 Days | 30 Days | Compliant | Not Compliant | Compliant | Compliant | Compliant | 1 | 0 | 1 | 1 | 1
LOGISTICS | LOG-02 | Packing Method | Export crate, waterproof | Export crate, waterproof | Cardboard box, waterproof liner | Plastic crate, stretch-wrapped | Export crate, waterproof | Export crate, waterproof | Compliant | Not Compliant | Partially Compliant | Compliant | Compliant | 1 | 0 | 0.5 | 1 | 1
DOCUMENTATION | DOC-01 | Manual Language | English & Russian | English & Russian | English | English & Russian | English & Russian | English & Russian | Compliant | Not Compliant | Compliant | Compliant | Compliant | 1 | 0 | 1 | 1 | 1
SUPPORT | SUP-01 | Warranty Period | ≥ 24 months | 24 Months | 18 Months | 12 Months | 24 Months | 36 Months | Compliant | Not Compliant | Not Compliant | Compliant | Compliant | 1 | 0 | 0 | 1 | 1
SUPPORT | SUP-02 | Local After-Sales Support | Not required but recommended | Available in Uzbekistan | Not available | Available | Available | Available | Compliant | Not Compliant | Compliant | Compliant | Compliant | 1 | 0 | 1 | 1 | 1
Payment | — | Payment Terms | 50% advance, 50% after final delivery | 30% advance, 70% after delivery | 50% advance, 50% after delivery | 20% advance, 80% after acceptance | Full payment after delivery | 40% advance, 60% after delivery | Compliant | Compliant | Compliant | Compliant | Compliant | 1 | 1 | 1 | 1 | 1
Payment | — | Cost per item | ≤ $500,000 | $3875 | $4,200 | $4,100 | $3,950 | $4,300 | Compliant | Not Compliant | Compliant | Compliant | Not Compliant | 1 | 0 | 1 | 1 | 0
[13 empty] | Overall compliance: | 94% | 42% | 64% | 78% | 78%