## Tender Metadata
Field | Value
Tender Title | Procurement of Medical Refrigeration Units for Immunization Program
Tender Ref No | MOH/PROC/2025/078
Tender Issuer | Ministry of Health, Uzbekistan
Submission Deadline | 2025-06-15
Contract Duration | 24 Months (including warranty)
Incoterms | DDP – Tashkent Regional Medical Depot
Delivery Timeline | Within 35 calendar days after contract signature
Payment Terms | 50% advance, 50% after final delivery
Evaluation Method | Most Economically Advantageous Tender (MEAT)
Required Quantity | 120 Refrigeration Units
Budget (estimated) | $500,000 USD
Currency | USD

## Technical_Specifications
Group | Spec Code | Parameter Name | Required Value | Unit | Mandatory | Weight (%) | Notes
PERFORMANCE | PER-01 | Temperature Range | +2 to +8 | °C | Yes | 10% | Must remain stable in ambient up to 43°C
PERFORMANCE | PER-02 | Temp Fluctuation | ≤ ±2 | °C | Yes | 5% | Should be measured over 24h test period
POWER | POW-01 | Power Supply | 220-240 | V AC | Yes | 3% | 50Hz, European socket compatible
POWER | POW-02 | Power Consumption | ≤ 1.5 | kWh/24h | Yes | 7% | Energy Star or equivalent certificate preferred
SAFETY | SAF-01 | Alarm Type | Visual + Audio | - | Yes | 5% | Should include over-temp and door open alarms
SAFETY | SAF-02 | Battery Backup Duration | ≥ 8 | Hours | Yes | 6% | For uninterrupted operation during outages
DESIGN | DES-01 | Gross Internal Capacity | ≥ 250 | Liters | Yes | 5% | Net usable storage for vaccines only
DESIGN | DES-02 | Internal Shelves | Adjustable, SS/plastic | - | Yes | 3% | Must be easy to sanitize
DESIGN | DES-03 | External Size (W×D×H) | ≤ 600×600×1800 | mm | No | 2% | Space is limited in storage rooms
CERTIFICATION | CER-01 | CE Marking | Required | - | Yes | 5% | For European compliance
CERTIFICATION | CER-02 | ISO 13485 | Manufacturer must have | - | Yes | 3% | Must be valid at time of bid
LOGISTICS | LOG-01 | Delivery Timeline | ≤ 35 | Days | Yes | 5% | After contract signature
LOGISTICS | LOG-02 | Packing Method | Export crate, waterproof | - | Yes | 2% | Heat-treated wooden crate
DOCUMENTATION | DOC-01 | Manual Language | English & Russian | - | Yes | 2% | Accompany each unit
SUPPORT | SUP-01 | Warranty Period | ≥ 24 | Months | Yes | 4% | Full parts and labor
SUPPORT | SUP-02 | Local After-Sales Support | Available in Uzbekistan | - | No | 1% | Bonus points for local support center